ENEL4CB – Computer engineering design project

Project title: The Vitality Watchdog System

Phase 1 report

Prepared by Orien S. Harichand (222012115)

Project focuses:

| Area/Field | Weight in project (%) |
| --- | --- |
| Biomedical engineering and Instrumentation | 30 |
| Computer Vision, Signal and Image processing | 30 |
| Embedded systems, IoT and edge computing | 10 |
| Artificial intelligence and machine learning | 10 |

GA assessment table:

| GA | ECSA definition | Implementation | Status |
| --- | --- | --- | --- |
| 1 - Problem solving | Identify, formulate, analyse, solve CEP | Researching and identifying a CEP that falls under my topic, while providing an acceptable solution and discussing other available options according to the literature sourced. | GREEN |
| 2 - Application of scientific and engineering knowledge | Apply knowledge of math, natural science, engineering fundamentals and an engineering speciality to solve CEP | Research into heart health metrics and photoplethysmography (PPG) signals; and how those metrics can be extracted from PPG signals via mathematical methods and implementation of said methods to extract heart rate. Research into the relationship between stress and the heart health metrics and how leveraging machine learning models to analyse metrics can improve accuracy of stress detection of stress states in a person. | BLACK |
| 3 - Engineering design | Perform creative, procedural and non-procedural design and synthesis of components, systems, engineering works, products and processes |  | BLACK |
| 6 - Professional and technical communication | Communicate effectively in oral and written mediums with engineering audiences and community | Creation, editing and review of design report documents to be submitted for phases 1, 2 and 3 of this module, along with demonstrations to be performed and oral presentations to be held. | BLACK |

# Abstract

Table of Contents

Abstract	2

1.	Introduction	3

2.	The Problem	4

2.1.	Problem identification	4

2.2.	Requirements and specifications	5

2.3.	Constraints	7

2.4.	Assumptions	7

3.	Literature Review	7

4.	The solution	10

4.1.	Proposed solution	10

4.2.	Alternative solutions	10

5.	Feasibility study	10

6.	Analysis and design of solution	11

6.1.	System design	11

6.2.	Subsystem design	11

6.2.1.	Subsystem 1: Signal Processing layer	11

6.2.2.	Subsystem 2: Data Analysis layer	11

6.2.3.	Subsystem 3: Alert system	12

6.3.	Target/Performance metrics	12

6.4.	Preliminary Testing	12

6.4.1.	Preliminary testing methods	12

6.4.2.	Preliminary testing results	12

7.	Overview of progress	12

8.	Conclusion	13

9.	References	13

# Introduction

People often don’t feel the severity of stress and cardiac issues and pass off some symptoms that could indicate cardiac issues (such as prolonged rapid/slow heart rates) as nothing to worry about. People experiencing severe and sudden cardiac issues may also have trouble calling for help and alerting others during the period of distress. Both of these issues can have severe impacts on a person’s health and may even lead to permanent damage should timely intervention not occur.

//put motiv here

//describe stakeholderas

This could be mitigated/lessened if the person had a way of automatically notifying themselves/others if they are in danger (such as when their heart rate (HR)/respiratory rate (RR) is too high/low) so that immediate action/measures can be taken to assist the person.

Currently, very few devices on the market have all of these capabilities in a singular system, and the ones that do fall into a cost range not suitable for the average household.

//conceptual block here

# The Problem

//scope of sys an

## Problem identification

There exists many devices that can measure heart health and detect stress levels in a person, such as the Google Fitbit Versa 4 (~R3000) [1], Garmin Vivosmart 5 (~R2500) [2], Wellue O2Ring (~R3000) [3] and Hume Band (~R4600), and others that feature subscription-based models to unlock the full benefits of the device – the Whoop band (~R6000/year for full benefits). These devices are problematic in the sense that they are far too costly for the average consumer who may require heart monitoring solutions. cheaper options exist, such as the Polar H10 (~R1800) [4] and Polar Verity Sense (~R1600)[5], but this range of alternatives provide only a select amount of features, such as dedicated heart rate monitoring sessions that need to be manually activated. Some of the solutions that exist are bulky (the Wellue O2Ring) and can inconvenience/discomfort the user during long-term wear. Furthermore, none of the devices identified had the ability to automatically alert others when periods of high stress were detected – they simply observed and informed the user via a watch interface/mobile application that they were indicating high stress levels when.

This highlights the need for a low-cost device that performs all the functions of higher-end alternatives (continuous heart rate monitoring in a non-invasive manner [6], stress detection and mobile interface) as well as possessing the ability to automatically alert the user and others when high stress states are detected.

Detecting heart and respiratory rates via non-invasive means was a challenge solved with the introduction of the electrocardiogram (ECG) [7] and photoplethysmography (PPG)[8]. However, real world factors affect these readings, such as motion artifacts [8], that can hamper accuracy of readings obtained. The accuracy can be improved upon by utilizing various types of filters[8], [9] and signal processing/preprocessing [8], but at the cost of power consumption experienced by the hardware and computational overhead taken for signal processing. This indicates the need for a balance between accuracy of results obtained, the hardware (sensors/filters etc.) that is used to implement the system and signal processing algorithms employed.

Detecting stress in a person can be achieved via a plethora of approaches. Some of the non-invasive approaches include image-based detection via digital cameras[10] and via wearable health devices [11]. This project focuses on a mobility-driven approach so as to allow stress monitoring that isn’t constrained to a singular room/place and will thus focus entirely on a wearable device approach.

Relating stress to heart health metrics can be achieved by analysing heart rate variability (HRV) [12] or Pulse rate variability (PRV) [13] metrics which are derived from ECG and PPG waveforms respectively. These metrics can be derived from the respective waveforms and analysed via various methods, which can be grouped into 2 categories – machine learning and algorithmic approaches[14], [15], [16], [17], [18], [19], [20] , with each approach and method offering different trade-offs. Machine learning approaches typically offer higher predictive accuracy, whereas algorithmic approaches provide deterministic, transparent calculations. Machine learning workflows require higher computational resources and training data, while algorithmic methods deliver low-latency and computationally efficient processing suitable for real-time edge devices, but remain more sensitive to motion artifacts and signal noise. This presents a further balance that is required between accuracy and computational efficiency of the system.

## Requirements and specifications

Below is a table of the requirements derived from the problem statement, complex problem identification and literature reviewed further down in the report. The project specifications are derived from the requirements.

|  | Requirements | Specifications with |
| --- | --- | --- |
| Functional | Continuous monitoring of vital signs. | Successful BLE pairing to hardware sensor |
|  |  | Implementation of rolling buffers to store data continuously. |
|  | Computation of heart rate, respiratory rate and HRV metrics. | Calculated heart rate accuracy of 65% before MA mitigation. |
|  | Mitigation of influences of Motion artifacts. | Signal pre-processing should have identification and filtering of motion artifacts prior to calculations. //put into sys design |
|  |  | Calculated heart rate accuracy of 70% after MA mitigation. |
|  | Evaluation of heart rate/HRV metrics to classify physiological stress states. | Accuracy of stress-state evaluation of 70%. |
|  |  | Identification of elevated sustained heart rates for 5 minutes. |
|  |  | Classification of stress states into unstressed, moderately stressed, severely stressed or extremely stressed for sustained heart rates. |
|  |  | Classification of HRV-based determination of stress into no stress, moderate stress or severe stress. |
|  | Alerting the user/others nearby of severely dangerous and extremely dangerous heart rates/stress states. | Successful UI update of heart rate and stress state. |
|  |  | Remote notifications are sent out to selected emergency contacts when extreme stress states are detected. |
|  |  | Successful audible alerts for 5 to 10s (seconds) upon detection of severe/extreme heart rates/stress. |
|  |  | A physical feedback (haptic feedback) device is triggered for 5 to 10s upon detection of severe/extreme heart rates/stress. |
|  |  | The system dispatches alerts within 60s from the moment extreme heart/stress states are determined. |
|  |  | The system dispatches alerts within 10s from the moment the “alarm window” lapses for severe stress/heart rates. |
| Non-functional | The full system should be cost-friendly. | The prototype of the system should cost less than |
|  | The system must be ergonomic for users. |  |

## Constraints

The system will NOT diagnose any specific diseases or health conditions; it will only determine general wellness and indicate stress levels of a person.

The system will focus on heart rate, heart rate variability and respiratory rate metrics only.

//don’t mention spo2 anywhere

## Assumptions

The user owns a modern smartphone that has:

- Bluetooth, WiFi and messaging capabilities (such as email and SMS)

- A WiFi connection or enough mobile data and/or airtime on their phone such that a message can be sent out from their phone successfully.

- A working audio speaker.

The user’s phone is kept within the broadcast range of the system’s sensor.

# Literature Review

Heart rate is the most commonly used physiological marker of stress in the literature[10], [11], [21], [22], [23], [24]. It rises reliably under both physical and psychological stressors, can be sensed cheaply and continuously and many studies report a clear increase in heart rate during a stressor compared to a resting baseline [10], [11], [22], [23], [25], [26]. This has been shown using stimuli such as (but not limited to) the Stroop test[17], mental arithmetic[14] and cold-water immersion[14]. Reported classification accuracies for stressed-versus-relaxed states, using only heart-rate features and simple classifiers such as AdaBoost, discriminant-function analysis, or SVM, range from roughly 83% to 92% [21]. This confirms heart rate as a workable primary signal.

The initial detection strategy proposed for this system establishes a baseline heart-rate pattern for each person, then flags stress as a percentage deviation from that baseline. A larger deviation indicates greater stress, and a top threshold triggers a "danger" state. This approach is supported in the literature – Sacha's [12] analysis of the link between heart rate and heart rate variability shows that both measures should be interpreted relative to a person's own resting physiology, not against a single population-wide value, because the same autonomic change produces different absolute shifts depending on a person's baseline heart rate. Using continuous data from a large number of wearable users, Presby et al. showed that resting heart rate and heart rate variability follow a predictable daily rhythm for each individual. By first correcting each user's readings for their own time-of-day baseline, the researchers were able to detect and measure the size of both physical and psychological stress events in everyday conditions. Heart rate stayed measurably raised for up to 30 minutes after a high-stress work period, and for several hours after vigorous exercise [22]. Natarajan et al.'s analysis likewise shows that normal heart-rate and heart-rate-variability values shift with age, sex, and time of day. This supports the case for a baseline that is personal to each user, rather than one fixed cut-off applied to everyone [27].

In terms of what percentage deviation to use, a clinical study that continuously monitored hospitalised patients found that a change of roughly 20% in heart rate or breathing rate, sustained for at least 20 minutes, worked well as an “alarm window rule”[28]. It reliably picked up genuine physiological decline while ignoring short-lived noise [28]. Hospital early-warning systems such as NEWS2 give their highest deterioration score to any heart rate at or below 40 beats per minute (bpm) or at or above 131, regardless of the patient's own normal range, because at those extremes the absolute rate is dangerous on its own [29]. Therefore, the system will organise the “danger” tiers as follows: heart rates deviating within 0-14% from the baseline as “safe heart rate/unstressed”, heart rates deviating within 15-20% from the baseline as “moderately dangerous heart rate/moderately stressed”, deviations greater than 20% as “severely dangerous heart rate/severely stressed” and heart rates lower than 40bpm and greater than 131bpm as “extremely dangerous heart rate/extremely stressed”. Due to the system focusing on deployment in real-world environments (non-clinical/hospital/lab-controlled conditions), the “alarm window rule” will be shortened to 5 minutes.

Heart rate has traditionally been measured using the electrocardiogram (ECG), still the clinical reference standard for the heart's electrical activity [7]. Its wired, multi-electrode setup, however, is not practical for a device that is worn continuously[11]. Photoplethysmography (PPG) is the usual wearable alternative. It works by shining light into the skin and sensing the resulting change in reflected light caused by blood-volume changes with each heartbeat, following the Beer–Lambert law of light absorption [8]. Importantly for this project, Tsai et al. showed that pulse intervals measured with PPG correlate very strongly with the intervals measured at the same time using ECG R2=(0.9837). This directly supports using PPG instead of ECG for both heart-rate sensing and for computing heart rate variability afterwards [17]. Other sensing methods exist, such as ballistocardiography and phonocardiography, which sense the heart's mechanical movement or its sounds rather than its electrical or optical signal[30]. These require more complex hardware and are less developed for low-cost consumer wearables than PPG [6], which supports the choice of a PPG-based sensor for this system.

Heart rate variability refers to the small fluctuations in the time between heartbeats, and it is a well-established, non-invasive marker of autonomic nervous system activity[31]. It is usually analysed using time-domain measures such as the standard deviation of normal-to-normal intervals (SDNN), the root mean square of successive differences (RMSSD), and the percentage of successive normal-to-normal intervals differing by more than 50 milliseconds (pNN50); geometric measures taken from the histogram of RR intervals, such as the stress index; frequency-domain measures such as high-frequency (HF), low-frequency (LF), and very-low-frequency (VLF) power and the LF/HF ratio; and nonlinear measures such as Poincaré-plot descriptors, entropy, and fractal dimension [31].

In terms of method, the literature splits stress-analysis approaches into rule-based (algorithmic) and machine-learning approaches, and both appear across heart rate, heart-rate-variability, and breathing-rate studies. On the rule-based side, Luo et al.[18] built a purely algorithmic, weighted composite index of heart-rate-variability parameters on a wearable ECG platform. Each parameter was weighted by its known influence on stress, so no single noisy parameter could dominate the score, while keeping the computation light enough for basic wearable hardware[18]. This is a genuine fallback option if a trained classifier turns out to be too demanding for the target hardware. On the machine-learning side, Tsai et al. [17] trained a decision-tree classifier on a 34-parameter heart-rate-variability feature set drawn from the Koldijk et al. SWELL dataset[32], and validated it against psychological questionnaire scores. The model reached 99.96% classification accuracy under lab conditions, with precision, recall, and F1 all around 0.999, and was compared directly against several alternatives on related tasks: a support vector machine (79.54%), a decision tree using voice and skin-conductance data instead of HRV (92.60%), a random forest (90.01%), and a convolutional neural network (94.8%) [17]. Tsai et al.’s classifier organises stress states either no stress, moderate stress or severe stress.

All of these alternatives’ accuracies were obtained on carefully curated datasets rather than real-world wearable data. PPG is sensitive to physical movement which makes it highly vulnerable to motion artifacts in wearable applications [8]. These motion artifacts corrupt the optical signal and introduce missing data, which can lead to estimation errors in both calculated heart rate and ultra-short-term heart rate variability metrics[33]. This is why a more conservative accuracy target is chosen as 70% after mitigation of MA’s and 65% before mitigation is implemented. It is also why, of the two approaches mentioned prior, the decision tree is a better consideration for this project specifically: it combines the lightest computational complexity of the machine-learning options tested with accuracy that matches or exceeds heavier alternatives.

Respiration and heart rate variability are closely linked[16]. Hernando et al. found that re-centring the high-frequency band on the person's actual breathing rate, rather than using a fixed 0.15–0.4Hz range, gives a more reliable read on sympathetic activation during stress. In their study, breathing rate on its own was the single strongest stress marker, with an area under the curve of 88.2% [16]. So while the decision tree can run on PPG-derived heart-rate variability alone, adding breathing rate as an extra feature (which can be extracted from the same PPG signal[34], [35]) is a well-supported way to improve accuracy later.

The specification that a local alert be sent within 60s (seconds) of extreme stress identification fits with the minimum window lengths shown to give reliable ultra-short-term heart-rate-variability estimates[20]: heart rate itself can be reliably estimated within about 10s, and the full set of frequency-domain features, including the LF/HF ratio, within about 50s. This leaves a workable margin of 10s for computing features, classifying stress, and sending the alert. For severe stress, the same 10s margin is used to trigger alerts after a heart rate that differs by 20% from the baseline is sustained during the “alarm window”.

The choice to trigger 5–10 seconds of vibration feedback also fits with what is known about how reliably people notice such alerts, even under high mental or physical load - a study on driver alerting found that a vibration lasting longer than about 3 seconds was accurately noticed under all tested conditions, including high-stress driving, while shorter or paused pulses were often misread [36]. The same reasoning is used for the 5-10s of audible alerts.

# The solution

## Proposed solution

## Alternative solutions

//include conceptual/detailed diagrams/flowcharts

//talk about ecg, app, alert mechanisms

## Feasibility study of proposed solution

//tools and components needed

# Analysis and design of solution

## System design

//maybe talk about decision tree model here

//flowchart of overall sys design

## Subsystem design

## Subsystem 1: Signal Processing layer

This subsystem is intended to provide signal processing services that converts raw PPG signals from a hardware sensor into usable data (Heart rate[8], respiratory rate [8], [35] and heart rate variability metrics[8], [15], [31]) to be interpreted by the data analysis layer. This layer will also handle pre-processing of the signal which includes identifying and filtering out motion artifacts[8], [33] to enhance signal quality.

Heart rate itself is derived by analysing the PPG signal to determine

Respiratory rates are determined by

Multiple heart rate variability metrics can be obtained from a PPG signal as discussed in the literature review. For this project, a decision tree model is implemented (detailed in subsystem 2 below) to analyse HRV parameters and determine stress [17]. This model requires only the following 34 HRV parameters for its analysis:

Time-domain parameters:

Relative time-domain parameters:

Frequency-domain parameters:

Nonlinear parameters:

//need equations and explanations for each parameter and each parameter type.

## Subsystem 2: Data Analysis layer

This layer provides analytical services to be used by the danger alert system. It takes in the values derived (heart rate, respiratory rate, rMSSD..) from the signal produced by the signal processing layer and utilizes a lightweight machine learning model to analyse them. A decision tree was chosen as the machine learning model as it provided high accuracies (99.96%) [17] while remaining lightweight. This project will attempt to achieve an accuracy of 70% due signal processing being done from first principles.

//put stuff about the model here..go into depth..try for alg first

## Subsystem 3: Alert system

This subsystem deals with human-machine-interaction. It provides the ability to inform the user of their stress state that is determined by the data analysis layer. It fulfils its functions by first notifying the user (via push notifications from the mobile application) of their stress state when they are determined to be moderately stressed. When they are detected to be severely stressed, however, this subsystem will provide physical and auditory feedback to the user. Auditory feedback is given through the users device’s phone speakers and physical feedback is provided by a haptic feedback cuff that incorporates a vibration motor, a coin cell battery to provide power to the circuit and a microcontroller to detect the incoming signal and toggle the vibration motor.

The circuit diagram of this implementation as well as the cuff design can be seen in the figures below:

//insert figures

//compare different motors and components and say why choose

## Target/Performance metrics

## Preliminary Testing

## Preliminary testing methods

## Preliminary testing results

## Feasibility of system

# Overview of progress

//gaant chart/timeline

The project is on track with 50% of all specifications being met. The mobile application successfully builds on an android device, pairs with a heart rate sensor and obtains data from it with accuracies and errors within acceptable tolerances. Further work is needed on the stress detection model’s implementation and, signal pre-processing and local notification/alert generation to achieve 75% of the project specs. To meet all project specs, future work should focus on implementing physical feedback to the alerts, remote notifications, implementing the app on devices running iOS and refining the analysis of data to cater to multiple age ranges, not to the averages currently used.

# Conclusion

# References

[1]	“Google Fitbit Versa 4.” Accessed: Jul. 26, 2026. [Online]. Available: https://store.google.com/us/product/fitbit_versa_4?hl=en-US&pli=1

[2]	Garmin and G. L. or its subsidiaries, “vívosmart® 5,” Garmin. Accessed: Jul. 26, 2026. [Online]. Available: https://www.garmin.com/en-US/p/782585/pn/010-02645-00/

[3]	“Wellue O2Ring Smart Ring Pulse Oximeter. Overnight Tracking.,” Wellue. Accessed: Jul. 06, 2026. [Online]. Available: https://getwellue.com/pages/o2ring-oxygen-monitor

[4]	“Polar H10 | Polar USA.” Accessed: Jul. 23, 2026. [Online]. Available: https://www.polar.com/us-en/sensors/h10-heart-rate-sensor

[5]	“Polar Verity Sense | Polar USA.” Accessed: Jul. 23, 2026. [Online]. Available: https://www.polar.com/us-en/products/accessories/polar-verity-sense

[6]	Q. Yong, L. Yang, A. Kardos, and Y. Zhao, “Non-invasive cardiovascular and vital signs monitoring techniques: review, challenges, and perspectives,” Measurement, vol. 258, p. 119472, Jan. 2026, doi: 10.1016/j.measurement.2025.119472.

[7]	“Electrocardiogram (ECG),” nhs.uk. Accessed: Jul. 23, 2026. [Online]. Available: https://www.nhs.uk/tests-and-treatments/electrocardiogram/

[8]	J. Park, H. S. Seok, S.-S. Kim, and H. Shin, “Photoplethysmogram Analysis and Applications: An Integrative Review,” Front. Physiol., vol. 12, Mar. 2022, doi: 10.3389/fphys.2021.808451.

[9]	D. Wu, Guan-Zheng Liu, M. Y. M. Wong, and Y.-T. Zhang, “The accuracy of respiratory rate estimation using electrocardiography and photoplethysmography,” in Proceedings of the 10th IEEE International Conference on Information Technology and Applications in Biomedicine, Corfu, Greece: IEEE, Nov. 2010, pp. 1–3. doi: 10.1109/ITAB.2010.5687633.

[10]	N. Molinaro et al., “Contactless Vital Signs Monitoring From Videos Recorded With Digital Cameras: An Overview,” Front. Physiol., vol. 13, p. 801709, Feb. 2022, doi: 10.3389/fphys.2022.801709.

[11]	D. Dias and J. Paulo Silva Cunha, “Wearable Health Devices—Vital Sign Monitoring, Systems and Technologies,” Sensors, vol. 18, no. 8, p. 2414, Jul. 2018, doi: 10.3390/s18082414.

[12]	J. Sacha, “Interaction between Heart Rate and Heart Rate Variability,” Noninvasive Electrocardiol, vol. 19, no. 3, pp. 207–216, May 2014, doi: 10.1111/anec.12148.

[13]	J. S. Burma et al., “Heart Rate Variability and Pulse Rate Variability: Do Anatomical Location and Sampling Rate Matter?,” Sensors, vol. 24, no. 7, p. 2048, Mar. 2024, doi: 10.3390/s24072048.

[14]	B. Villatte, S. A. D. Kizuk, J.-M. Lina, A. Vinet, and S. Hébert, “Heart rate variability (HRV) during acute stress: a comparison of three methods for time–frequency analysis,” Physiol. Meas., vol. 47, no. 3, p. 035001, Mar. 2026, doi: 10.1088/1361-6579/ae3ec7.

[15]	T. Pereira, P. R. Almeida, J. P. S. Cunha, and A. Aguiar, “Heart rate variability metrics for fine-grained stress level assessment,” Computer Methods and Programs in Biomedicine, vol. 148, pp. 71–80, Sep. 2017, doi: 10.1016/j.cmpb.2017.06.018.

[16]	A. Hernando et al., “Inclusion of Respiratory Frequency Information in Heart Rate Variability Analysis for Stress Assessment,” IEEE J. Biomed. Health Inform., vol. 20, no. 4, pp. 1016–1025, Jul. 2016, doi: 10.1109/JBHI.2016.2553578.

[17]	Y.-Y. Tsai, Y.-J. Chen, Y.-F. Lin, F.-C. Hsiao, C.-H. Hsu, and L.-D. Liao, “Photoplethysmography-based HRV analysis and machine learning for real-time stress quantification in mental health applications,” APL Bioengineering, vol. 9, no. 2, p. 026103, Jun. 2025, doi: 10.1063/5.0256590.

[18]	J. Luo et al., “Quantitative analysis of heart rate variability parameter and mental stress index,” Front. Cardiovasc. Med., vol. 9, p. 930745, Jul. 2022, doi: 10.3389/fcvm.2022.930745.

[19]	W. Von Rosenberg, T. Chanwimalueang, T. Adjei, U. Jaffer, V. Goverdovsky, and D. P. Mandic, “Resolving Ambiguities in the LF/HF Ratio: LF-HF Scatter Plots for the Categorization of Mental and Physical Stress from HRV,” Front. Physiol., vol. 8, p. 360, Jun. 2017, doi: 10.3389/fphys.2017.00360.

[20]	L. Salahuddin, J. Cho, M. G. Jeong, and D. Kim, “Ultra Short Term Analysis of Heart Rate Variability for Monitoring Mental Stress in Mobile Settings,” in 2007 29th Annual International Conference of the IEEE Engineering in Medicine and Biology Society, Lyon, France: IEEE, Aug. 2007, pp. 4656–4659. doi: 10.1109/IEMBS.2007.4353378.

[21]	A. Arsalan, M. Majid, I. F. Nizami, W. Manzoor, S. M. Anwar, and J. Ryu, “Human Stress Assessment: A Comprehensive Review of Methods Using Wearable Sensors and Non-wearable Techniques,” 2022, arXiv. doi: 10.48550/ARXIV.2202.03033.

[22]	D. M. Presby, S. R. Jasinski, and E. R. Capodilupo, “Wearable derived cardiovascular responses to stressors in free-living conditions,” PLoS ONE, vol. 18, no. 6, p. e0285332, Jun. 2023, doi: 10.1371/journal.pone.0285332.

[23]	R. Chen, “Wearable Medical Devices: Regulatory Pathway Guide (FDA, EU MDR 2026).” Accessed: Jul. 17, 2026. [Online]. Available: https://meddeviceguide.com/blog/wearable-medical-devices-regulatory-pathway-guide

[24]	A. Aliverti, “Wearable technology: role in respiratory health and disease,” Breathe, vol. 13, no. 2, pp. e27–e36, Jun. 2017, doi: 10.1183/20734735.008417.

[25]	M. Harford, J. Catherall, S. Gerry, J. Young, and P. Watkinson, “Availability and performance of image-based, non-contact methods of monitoring heart rate, blood pressure, respiratory rate, and oxygen saturation: a systematic review,” Physiol. Meas., vol. 40, no. 6, p. 06TR01, Jun. 2019, doi: 10.1088/1361-6579/ab1f1d.

[26]	J. E. Sharman et al., “Automated ‘oscillometric’ blood pressure measuring devices: how they work and what they measure,” J Hum Hypertens, vol. 37, no. 2, pp. 93–100, Feb. 2023, doi: 10.1038/s41371-022-00693-x.

[27]	A. Natarajan, A. Pantelopoulos, H. Emir-Farinas, and P. Natarajan, “Heart rate variability with photoplethysmography in 8 million individuals: a cross-sectional study,” The Lancet Digital Health, vol. 2, no. 12, pp. e650–e657, Dec. 2020, doi: 10.1016/S2589-7500(20)30246-6.

[28]	G. M. Peters, R. V. Peelen, V. J. Gilissen, M. V. Koning, W. H. Van Harten, and C. J. M. Doggen, “Detecting Patient Deterioration Early Using Continuous Heart rate and Respiratory rate Measurements in Hospitalized COVID-19 Patients,” J Med Syst, vol. 47, no. 1, p. 12, Jan. 2023, doi: 10.1007/s10916-022-01898-w.

[29]	Royal College of Physicians, “National Early Warning Score (NEWS) 2: Standardising the assessment of acute-illness severity in the NHS,” Royal COllege of Physicians. [Online]. Available: 10.1007/s10916-022-01898-w

[30]	J. Nedoma et al., “Comparison of BCG, PCG and ECG signals in application of heart rate monitoring of the human body,” in 2017 40th International Conference on Telecommunications and Signal Processing (TSP), Jul. 2017, pp. 420–424. doi: 10.1109/TSP.2017.8076019.

[31]	Roman М. Baevsky and Anna G. Chernikova, “Heart rate variability analysis: physiological foundations and main methods,” Cardiometry, no. 10, pp. 66–76, May 2017, doi: 10.12710/cardiometry.2017.10.6676.

[32]	S. Koldijk, M. Sappelli, S. Verberne, M. A. Neerincx, and W. Kraaij, “The SWELL Knowledge Work Dataset for Stress and User Modeling Research,” in Proceedings of the 16th International Conference on Multimodal Interaction, Istanbul Turkey: ACM, Nov. 2014, pp. 291–298. doi: 10.1145/2663204.2663257.

[33]	A. Rossi, D. Pedreschi, D. A. Clifton, and D. Morelli, “Error Estimation of Ultra-Short Heart Rate Variability Parameters: Effect of Missing Data Caused by Motion Artifacts,” Sensors, vol. 20, no. 24, p. 7122, Dec. 2020, doi: 10.3390/s20247122.

[34]	I. Costanzo, D. Sen, L. Rhein, and U. Guler, “Respiratory Monitoring: Current State of the Art and Future Roads,” IEEE Rev. Biomed. Eng., vol. 15, pp. 103–121, 2022, doi: 10.1109/RBME.2020.3036330.

[35]	J. Lázaro et al., “Pulse photoplethysmography derived respiration for obstructive sleep apnea detection,” in 2017 Computing in Cardiology (CinC), Sep. 2017, pp. 1–4. doi: 10.22489/CinC.2017.273-229.

[36]	N. J. Harrington and C. Narayanaswami, “Method and system for improving driver safety and situational awareness,” US7692552B2, Apr. 06, 2010 Accessed: Sep. 12, 2026. [Online]. Available: https://patents.google.com/patent/US7692552B2/en